## Uncertanity Evaluation
ESTIMATION OF UNCERTAINTY
PLANT - 2
1. Laboratory : |  | SANSERA ENGINEERING LTD. |  |  |  | Date:
2.Device under calibration:
Range : |  | 0-10 / 0.010mm
Resolution: |  | NA
3.Standards / Equipment used for calibration :
Sl.No. | Standard / Equipment | Range mm |  | Resolution mm |  | Uncertainty µm | Traceability |  | Uncertainty
1 | ULM Machine | 0-600 |  | 0.001µm |  | 0.056 | Nabl Lab No:CC-3792 and Vide certificate no: RMTL/01/324101569-A1 Calibrated On: 13.05.2024 valid upto 13.05.2026 |  | 0.08+0.82(L)
[9 empty] | 0.08 | 0.0164
0.0964
4. Temperature during calibration (Deg C) :
| Beginning | Middle |  | End |  | Mean | RH%
46.3
47.1
5. Type A | [7 empty] | 48.9
| Number of Readings = |  |  | 10 |  |  |  | 47.4333
| Readings mm | Mean | µm |  |  | Standard Deviation
| 0 | 0.0014 | -1.4 |  | 1.96
| 0 |  | -1.4 |  | 1.96
| 0 |  | -1.4 |  | 1.96
| 0.002 |  | 0.6 |  | 0.36
| 0 |  | -1.4 |  | 1.96 | 1.4298
| 0.002 |  | 0.6 |  | 0.36
| 0.001 |  | -0.4 |  | 0.16
| 0.003 |  | 1.6 |  | 2.56
| 0.004 |  | 2.6 |  | 6.76
| 0.002 |  | 0.6 |  | 0.36
18.4
6.0 Standard deviation of the mean : |  |  |  | n = |  | 0.45216 | µm
7.0 Degree of Freedom : |  |  |  | 9
Study Conducted by: Uday. Witnessed By :

## 2nd Page
ESTIMATION OF UNCERTAINTY
PLANT - 2
8.0 Type B Evaluation
Source of Uncertainity |  |  |  | Sensitivity Co-Efficient
U1 = | 0.056 | µm | Uncertainty reported in the calibration certifiacte of ULM | 1
U2= | 0.11 | µm | Uncertainty Reported due to the Accuracy of the ULM | 1
U3 = | 5 | µm | Uncertanity due to resolution of Dial indicator | 1
U4 = | 0.3 | °C | Uncertanity mentioned in the calibration Report of Thermometer (0.020*0.3*(11.5+11.5)/2=0.02µm/°C | 0.07 | U5 | 0.069
U5 = | 0.21 | °C | Uncertanity due to error of the thermometer (0.020*0.21*(11.5+11.5)/2=0.07µm/°C | 0.05 | U6 | 0.0483
U6 = | 0.3 | °C | Uncertainty due to the difference in temperature of calibration & 20 deg centigrate. | 0.23 | U7 | 0.23
Ci = alphaavg* length = (11.5+11.5)/2 *0.020 = 0.23
U7= | 0.5 | °C | Uncertainty due to temperature difference in Snap Gauge & reference item | 0.23 | U8 | 0.23
Alphaavg*Length = (11.5+4.7)/2mm/m/°C * 0.020 = 0.23mm/°C
U8 = | 0.5 | °C | Uncertainty due to temperature variation during calibration | 0.23 | U9 | 0.23
Alphaavg* length = (11.5+11.5)/2 *0.02 = 0.23
U9 = | 1.62 | µ/°C/mtr | Uncertainty due to thermal expansion coefficient (10% of root mean square of alpha) | 0.04 | U10 | 0.04
Environmental Control* Length = 2.0°C * 0.020m = 0.04m°C
U10 = | 1.4298 | µm | Type A | 1
132.25
Study Conducted by: Uday. Witnessed By :

## 3rd page
ESTIMATION OF UNCERTAINTY
PLANT - 2
UNCERTAINTY BUDGET :
Source of Uncertanity X i | Estimate x i | Limits ± xi (µm) | Probability Disturbution Type A or TypeB |  | Standard uncertainity u(xi) |  | Sensitivity Co-efficient ci |  | Uncertainity Contribution ui (y) in µm |  | Degree of freedom V i
U1 | 0.056 | 0.028 | Normal-B | 2 | 0.028 | µm | 1 | µm | 0.028 | 0.000784 | ∞
U2 | 0.11 | 0.055 | Rectangular-B | 3 | 0.063509 | µm | 1 | µm | 0.063509 | 0.0040333 | ∞
U3 | 5 | 2.5 | Rectangular-B | 3 | 2.8868 | µm | 1 | µm | 2.8868 | 8.3333 | ∞
U4 | 0.3 | 0.15 | Normal-B | 2 | 0.15 | °C | 0.07 | µm/°C | 0.0105 | 0.00011025 | ∞
U5 | 0.21 | 0.105 | Rectangular-B | 3 | 0.12124 | °C | 0.05 | µm/°C | 0.0060622 | 3.675e-05 | ∞
U6 | 0.3 | 0.15 | Rectangular-B | 3 | 0.17321 | °C | 0.23 | µm/°C | 0.039837 | 0.001587 | ∞ | 5
U7 | 0.5 | 0.25 | Rectangular-B | 3 | 0.28868 | °C | 0.23 | µm/°C | 0.066395 | 0.0044083 | ∞
U8 | 0.5 | 0.25 | Rectangular-B | 3 | 0.28868 | °C | 0.23 | µm/°C | 0.066395 | 0.0044083 | ∞
U9 | 1.62 | 0.81 | Rectangular-B | 3 | 0.93531 | µ/°C/mtr | 0.04 | µm/°C | 0.037412 | 0.0013997 | ∞
U10 | 1.4298 | 0.71492 | Type-A | 5 | 0.45216 | µm | 1 | µm | 0.45216 | 0.20444 | 9
8.5545
| Combined Standard Uncertainty = |  |  | 2.9248 |  |  | mm
| Effective Degree of Freedom = |  |  | 15757.4435
(Calculated Value)
| Coverage Factor = |  |  | 2
| Expanded Uncertainty = |  |  | 5.8496 |  |  | mm
Study Conducted by: Uday. Witnessed By :

## Certificate
Calibration Results ( Values are in µm)
| 0 | 100 | 200 | 300 | 400 | 500 | 600 | 700 | 800 | 900
Turn - 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Turn - 2 | 0 | 0 | -10 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Turn - 3 | -10 | 0 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | 0
Turn - 4 | 0 | 0 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | 0
Turn - 5 | 0 | 0 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | 0
Turn - 6 | 0 | 0 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | -10
Turn - 7 | -10 | -10 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | 0
Turn - 8 | 0 | -10 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | -10
Turn - 9 | 0 | 0 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | 0
Turn - 10 | 0 | 0 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | 0
Turn - 11 | -10 | -10 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | -10
Turn - 12 | -10 | -10 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | -10
Allowable Error for L.C.- 0.01 mm Dial Gauge as per IS 2092-1983 (Values are in mm)
Parameter |  |  |  |  | Permissible Error |  |  | Observed Error
Hisyteris |  |  |  |  | 0.003 |  |  | 0
Sensitivity |  |  |  |  | 0.003 |  |  | 0
Repeatability |  |  |  |  | 0.002 |  |  | 0.001
Any 1/10th rev. |  |  |  |  | 0.005 |  |  | 0.001
Any 1 rev. |  |  |  |  | 0.01 |  |  | 0.002
Any 2 revs. |  |  |  |  | 0.015 |  |  | 0.001
Any large interval up to 10revs. |  |  |  |  | 0.02 |  |  | 0.002
EXPANDED UNCERTAINTY : ±6.5µm ( The uncertainty of measurement is expressed at95.45% Confidence with coverage factor K-2)

## Observation
Upward Direction
| 0 | 100 | 200 | 300 | 400 | 500 | 600 | 700 | 800 | 900
Turn - 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Turn - 2 | 0 | 0 | -10 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Turn - 3 | -10 | 0 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | 0
Turn - 4 | 0 | 0 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | 0
Turn - 5 | 0 | 0 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | 0
Turn - 6 | 0 | 0 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | -10
Turn - 7 | -10 | -10 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | 0
Turn - 8 | 0 | -10 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | -10
Turn - 9 | 0 | 0 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | 0
Turn - 10 | 0 | 0 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | 0
Turn - 11 | -10 | -10 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | -10
Turn - 12 | -10 | -10 | ~ | ~ | ~ | ~ | ~ | ~ | ~ | -10
Hisyteris: |  |  | Sensitivity |  |  |  | Repeatability: |  | 0.001
Type - A (Repeatability)
Trail-1 | Trail-2 | Trail-3 | Trail-4 | Trail-5
0.001 | 0.001 | 0.001 | 0.001 | 0.001
Any 1/10th Revision |  |  |  | 0.001
Any 1 Revision |  |  |  | 0.002
Any 2 Revision |  |  |  | 0.001
Any large interval up to 10 Revision |  |  |  | 0.002